## HSF EO Comb
| Keyword | Value | Required Authorization to Write to Keyword | Read/Write | Type | Description | Additional Notes | Unit | Acceptable Range |  | Tolerance | Default Value | Comb Python Module Function Name
[9 empty] | Min | Max
LFC High Level Control (Simple)
| LFC_HK_ATT |  | N | R/W | float | Comb Attenuation H and K band (Intensity) | Controls the comb variable optical attenuator to adjust amount of attenuation on the H and K band comb light. Should be calibrated to match to stellar target intensity. Default value should be maximum attenuation to avoid overexposure of spectrograph detector when comb light fist accessed by user. 0-60 dB attenuation | dB | 0.01 | 60 | 0.01 | 60
| LFC_FLT_ATT |  | N | R/W | Integer (? Need check) | Control position of the slider holding the neutral density filters in the spectral flattener | ADD later Read manual of the slider from Flattener (Thorlab ELL 12) Note this is controlled on the Menlo Flattener laptop presently. | dB | TBD | TBD | TBD | TBD
| LFC_PMP_ATT |  | N | R/W | float | Pump Attenuation (Intensity) | Controls the comb pump line variable optical attenuator to adjust amount of attenuation on the combpump line. Acts as a wavelength flag near 1560 nm. Default value should be maximum attenuation to avoid overexposure of spectrograph detector when comb light fist accessed by user. | dB | 0.01 | 60 | 0.01 | 60 |  | I
Need to break up? | LFC_YJ_ON_OFF |  | N | R/W | Boolean | Open or close the shutter of flattened YJ comb output | Enables comb light to be sent from spectral flattener by opening the shutter inside the unit. This is the keyword to be accessed normally by an operator. +5V TTL signal to open; default value of 0 leaves shutter closed. 0: Closed. 1: Open |  | TTL |  |  | 0
Need to break up? | LFC_HK_ON_OFF |  | N | R/W | Boolean | Open or close the shutter after the HK flattener output | Enables comb light to be sent from spectral flattener by opening the shutter inside the unit. This is the keyword to be accessed normally by an operator. +5V TTL signal to open; default value of 0 leaves shutter closed. 0: Closed. 1: Open. Note - this shutter is currently controlled on the Menlo laptop. |  | TTL |  |  | 0
| LFC_YJ_ATT |  | N | R/W | float | Comb Attenuation Y and J band (Intensity) coming from KPF comb (Nem) | Controls the comb variable optical attenuator to adjust amount of attenuation on the Y and J band comb light from KPF comb. Should be calibrated to match to stellar target intensity. Default value should be maximum attenuation to avoid overexposure of spectrograph detector when comb light fist accessed by user. 0-60 dB attenuation | dB | 0.01 | 60 | 0.01 | 60
New Item | LFC_YJ_HK |  | N | R/W | String | Selects either HK or yJ comb to be sent to specctrograph | the NIR Astrocomb provides the HK comb light, and the KPF comb provides the yJ light; both signals are connected to a 2X2 optical switch and the user can select between the two choices. This KTL keyword sets the selected 2X2 switch channel.
Need to combine? | LFC_DOWN | lockout keyword | Y | R/W* | Boolean | Comb off | Full system power down. What can report back the status is off? |  | TTL
| LFC_UP | lockout keyword | Y | R/W* | Boolean | Comb startup | Restarts the comb from a full-off state; System should report "starting up" while going through startup procedure/algorithm |  | TTL
| LFC_STATUS |  | N | R | String | Reports the status of overall comb status. | Will be 'READY'¦'WARMING_UP'¦'OFF'¦'FAULT_[code]'¦'SAFE_MODE'; Safe mode means only mini comb running, high power EDFA is off.
See which of these parameters are regurally wanted and should be assigned a separate keyword | LFC_COMP |  | N | R | String | Report the status of every component inside the rackmount. | This will be a slow command, it queries each component in the comb and returns all status. Will be useful for trouble shooting.
| LFC_REFLT |  | Y | R/W |  | Reflatten comb (single event) | 1. TBD – will be used if have SLM for spectral flattening; (2-7-22_2. reflattening will be automatic if Menlo flattener purchased at a refresh rate internal to the flattener software; If Nem's DSF PIC based flattener is used, will also have a refresh rate - see if this is something that is user-variable. | [5 empty] | Don't have Menlo's code yet
| LFC_FLT_RTE |  | Y | R/W |  | Reflattener refresh rate | See "Reflatten comb" comments 2 and 3. | Hz |  |  |  |  | Don't have Menlo's code yet
See which of these parameters are regurally wanted and should be assigned a separate keyword | LFC_FLT_SPC |  |  | R |  | Get spectrum from Arcopitx spectrum in flattener
LFC Low Level Control and Operation
| LFC_RIO_T | lockout keyword | Y | R/W |  | RIO pump laser temperature | Typically control temperature, monitor current | C | -5 | 75 |  | 25
| LFC_RIO_I | lockout keyword | Y | R/W |  | RIO pump laser current | Typically control temperature, monitor current | mA | TBD | TBD |  | 145
| LFC_EDFA27_P | lockout keyword | Y | R/W |  | Small EDFA (500 mW) 1 output power | Use constant power mode (APC mode) ; Read is monitored, write is setpoint | mW | 0 | 630 |  | 630
| LFC_EDFA27_ONOFF | lockout keyword | Y | R/W | Boolean | Small EDFA (500 mW) 1 emission on/off
| LFC_EDFA13_P | lockout keyword | Y | R/W |  | Small EDFA (20 mW) 2 output power | Use constant power mode (APC mode) ; Read is monitored, write is setpoint | mW | 0 | 20 (TBD) |  | 20
| LFC_EDFA13_ONOFF | lockout keyword | Y | R/W | Boolean | Small EDFA (20 mW) 2 emission on/off
| LFC_EDFA23_P | lockout keyword | Y | R/W |  | Small EDFA (200 mW) 3 output power | Use constant power mode (APC mode) ; Read is monitored, write is setpoint | mW | 0 | 20 (TBD) |  | 20
| LFC_EDFA23_ONOFF | lockout keyword | Y | R/W | Boolean | Small EDFA (200 mW) 3 emission on/off
| LFC_RFAMP_I | lockout keyword | N | R |  | RF Amp current | Read only, not set | Amp |  |  | +/- 0.2 | 4.4(Typical)
| LFC_RFAMP_V | lockout keyword | Y | R/W |  | RF Amp voltage | This is the control value; Caution! no voltage ramp allowed - can only go directly from off to 30 V. | V | 30 | 30 |  | 30
| LFC_RFOSCI_I | lockout keyword | N | R |  | RF Oscillator input current | Read only, not set | mA | 0.35 | 0.7 |  | 0.41 (operating)/0.62(warming up)
| LFC_RFOSCI_V | lockout keyword | Y | R/W |  | RF Oscillator input voltage | Check whether manufacturer recommends ramp or not. | V | 15 | 15 |  | 15
| LFC_IM_BIAS | lockout keyword | Y | R/W | float | Mini-comb amplitude lock offset bias | Amplitude lock voltage bias, is set by the locking loop, applied by RP FPGA | V | -3 | 3 | 0.001
| LFC_IM_RF_ATT | lockout keyword | Y | R/W | float | Mini-comb RF voltage controlled variable attenuation setting |  | V | 0 | 10
| LFC_WSP_PHASE | lockout keyword | Y | R/W | float array shape=(2, N) | Waveshaper Dispersion Compensation | Controls the dispersion setting in the Finisar Waveshaper | ps |  |  |  | 2.1 ps/nm
| LFC_WSP_ATTEN | lockout keyword | Y | R/W | float array shape=(2, N)
| LFC_PTAMP_PRE_P | lockout keyword | N | R |  | High-power EDFA pre-amp output power | monitors output power from the Pritel pre-amp | mW | 50 | 140 |  | 135
| LFC_PTAMP_OUT | lockout keyword | N | R |  | High-power EDFA output power | Read only, not set | W | 0 | 4
| LFC_PTAMP_I | lockout keyword | Y | R/W |  | High-power EDFA pump current |  | Amp | 0 | 4.2 |  | 3.6
| LFC_PTAMP_ONOFF | lockout keyword | Y | R/W | Boolean | High-power EDFA emission on/standby
| LFC_PTAMP_LATCH | lockout keyword | Y | R/W | Boolean | RESET preamp and pwramp latching circuits
| LFC_WGD_T | lockout keyword | Y | R/W | float | Waveguide TEC control
| LFC_PPLN_T | lockout keyword | Y | R/W | float | PPLN TEC control
Not using the filter cavity | FC_T
Not using the filter cavity | FC_V
Not using the filter cavity | FC_PID_P
Not using the filter cavity | FC_PID_I
Not using the filter cavity | FC_PID_D
Not using the filter cavity | FC_PID_OFFSET
Not using the filter cavity | FC_PID_SETV
Not using the filter cavity | FC_PID_MEAV
Not using the filter cavity | FC_PID_ERR
Minicomb intensity servo | IM_PID_P
IM_PID_I
IM_PID_D
IM_PID_OFFSET
IM_PID_SETV
IM_PID_MEAV
IM_PID_ERR
Rubidium lock servo | RB_PID_P
RB_PID_I
RB_PID_D
RB_PID_OFFSET
RB_PID_SETV
RB_PID_MEAV
RB_PID_ERR
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle RIO laser
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle EDFA 1
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle EDFA 2
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle Waveshaper
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle Frequency Counter
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle Rubidium Clock
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle High Power EDFA
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle RF Oscillator
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle RF amplifier 1
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle RF amplifier 2
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle FPGA 1
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle FPGA 2
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle Flattener/Arcoptix spectrometer
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle Flattener/SLM
Duplicate |  | lockout keyword | Y | R/W | Boolean | WPS power cycle FPGA 1
Duplicate |  | lockout keyword | Y | R/W | Boolean | WPS power cycle FPGA 2
|  | lockout keyword | Y | R/W | Boolean | WPS power cycle FPGA 3
Don't need |  | lockout keyword | Y | R/W | Boolean | WPS power cycle Fiber laser comb pump diode
Don't need |  | lockout keyword | Y | R/W | Boolean | WPS power cycle Fiber laser comb amplifier diode 1
Don't need |  | lockout keyword | Y | R/W | Boolean | WPS power cycle Fiber laser comb amplifier diode 2
Don't need |  | lockout keyword | Y | R/W | Boolean | WPS power cycle Fiber laser comb FPGA
Don't need |  | lockout keyword | Y | R/W | Boolean | WPS power cycle Fiber laser comb piezo controller
New item |  | lockout keyword | Y | R/W | Boolean | WPS power cycle oscilloscope
New item |  | lockout keyword | Y | R/W | Boolean | WPS power cycle HK shutter
No door interlock |  | lockout keyword | N | R |  | Cabinet front door interlock | placeholder
No door interlock |  | lockout keyword | N | R |  | Cabinet back door interlock | placeholder
No enclosure interlock |  | lockout keyword | N | R |  | Optics encloser interlock | placeholder
| LFC_EMR |  | Y | R/W | Boolean | Minicomb only | Turns off the high power EDFA and only operates minicomb; no comb light will be output from the system; System should also report back this status
LFC Health and Status
| LFC_GPS_LOCKED |  | N | R | Boolean | GPS locked | Reporting from SRS rubidium clock.
| LFC_RF_COUNTER |  | N | R | Number | Repetition Rate (comb spacing) | Reporting from Pendulum frequency counter | Hz |  |  |  | 16000000000
| LFC_OSC_LOCKED |  | N | R | Boolean | Drive Oscillator Locked | Reporting from Nexyn oscillator through DAQ
| LFC_TIMESTAMP_LOCAL |  | N | R | Number | Comb Computer Time stamp | Should be synchronized with facility clock
Not needed? Rely instead on lock signal for rubdium servo and assume 1560.492 nm? Previously said: "NO FIBER COMB NOW. WILL RETURN OSA READ FROM DOUBLER INSTEAD" | LFC_RIO_WAVELENGTH |  | N | R | Number | RIO Pump center wavelength | Read from fiber comb, calibrated to offset read at wavemeter; can only measure beatnote from FPGA, not absolute wavelength; purchase wavemeter?
| LFC_RIO_LOCKED |  | N | R | String | Pump laser status (on, locked, standby) | See FPGA2. What is used to determine? Error signal? Standby means laser emission on but lock status indeterminate? Read beat note - see if SNR too low, if so, not locked.
| LFC_RACK_TEMP |  | N | R | Number Array | Comb board and assembly operating temperature (15+ signals) | Instrument rack top, center, bottom; comb board RF cooling inlet1, outlet1, inlet2, outlet2; fiber comb board baseplate inlet and outlet; comb generator baseplate temperature inlet, outlet; flattener box; See specific items rows 84-98 | C |  | 30
Redundant | ----- |  | N | R | String | State reporting (on, off, standby, starting up) | Summary of current comb system state (sending light to spectrograph or not, comb in process of startup or locking, or in error)
Ambiguous? ASK Steph. -MD comment | LFC_FLT_COUNTS |  | N | R | Number | Average counts | From spectrometer internal to waveshaper -The values will have to be measured/calibrated. The idea is to get an intensity reading as seen by the flattener spectrometer. Gives us an idea of comb intensity injected into the flattener since we have no other optical power reading. TBD if necessary, but could provide insight if performance is degrading over time. Perhaps better just to read out spectrum periodically? Redundant with LFC_FLT_SPEC? | counts | 0 | TBD
| LFC_FLT_STATUS |  | N | R | String | Comb flattener spectrometer status | See comment for "Average Counts". See what parameters are reported for the Arcoptix spectrometer
| LFC_FLT_SPEC |  | N | R | Array | Comb flattener spectrometer spectrum | instantaneous comb spectrum as seen on the Arcoptix spectrometer inside the flattener unit.
| LFC_FC_LOCKED |  | N | R | Boolean | Filter cavity locked (SRS)
| LFC_IM_LOCKED |  | N | R | Boolean | Minicomb amplitude locked (SRS)
Redundant | LFC_RIO_LOCKED |  | N | R | Boolean | Pump laser relative to Rb locked (SRS)
|  | Temperatures | NOTE. implement these as individual keywords, not array
| LFC_T_GLY_RACK_IN |  | N | R | Number | Glycol inlet temp to instrument rack
| LFC_T_GLY_RACK_OUT |  | N | R | Number | Glycol outlet temp from instrument rack
| LFC_T_GLY_EOCB_IN |  | N | R | Number | Glycol inlet temp to EO comb board
| LFC_T_GLY_EOCB_OUT |  | N | R | Number | Glycol outlet temp from EO comb board
| LFC_T_GLY_FLB_IN |  | N | R | Number | Glycol inlet temp to fiber laser comb board
| LFC_T_GLY_FLB_OUT |  | N | R | Number | Glycol outlet temp from fiber laser comb board
| LFC_T_GLY_RFAMP1_IN |  | N | R | Number | Glycol inlet temp to RF amplifier block 1
| LFC_T_GLY_RFAMP1_OUT |  | N | R | Number | Glycol outlet temp from RF amplifier block 1
| LFC_T_GLY_RFAMP2_IN |  | N | R | Number | Glycol inlet temp to RF amplifier block 2
| LFC_T_GLY_RFAMP2_OUT |  | N | R | Number | Glycol outlet temp from RF amplifier block 2
| LFC_T_RACK_TOP |  | N | R | Number | Instrument rack top temp
| LFC_T_RACK_MID |  | N | R | Number | Instrument rack middle temp
| LFC_T_RACK_BOT |  | N | R | Number | Instrument rack bottom temp
| LFC_T_FLAT |  | N | R | Number | Flattener temp
| LFC_T_FILT |  | N | R | Number | Filter cavity temperature
Fiber Laser Comb Reference
Not using |  |  | N | R |  | Frep |  | MHz |  |  |  | 156
Not using |  |  | N | R |  | Fceo |  | MHz
Not using |  |  | N | R | Boolean | Frep locked
Not using |  |  | N | R | Boolean | Fceo locked
Not using |  |  | Y | R/W |  | Slow Piezo voltage | Ask Mahmood | V | 0 | 80
Not using | [5 empty] | Fast Piezo voltage | Ask Mahmood
Not using |  |  | N | R |  | Fiber comb temperature
Not using |  |  | Y | R/W |  | Fiber comb oscillator TEC voltage
Not using |  |  | N | R |  | Fiber comb output spectrum | May not have OSA to read
Not using |  |  | Y | R/W |  | Fiber comb pump diode current |  | mA
Not using |  |  | Y | R/W |  | Fiber comb amplifier diode 1 current |  | mA
Not using |  |  | Y | R/W |  | Fiber comb amplifier diode 2 current |  | mA
Not using |  |  | N | R |  | Reference Oscillator (200 MHz) Frequency | Might not need? | MHz
Not using |  |  | N | R |  | Reference Oscillator (1 GHz) Frequency |  | MHz
LFC Communications
| LPT_RPT |  |  |  |  | Handshake to the laptop check if laptop is online
| LPT_RST |  |  |  |  | Reset from blue-screen death. Not sure this is possible yet: may require manual reset.
OSA fiber switches
OSA_WLSPAN
OSA_WLCENT
OSA_SENS
OSA_RES
OSA_REF_P
OSA_RUN
Not using | OSA_SOURCE |  |  |  |  | Select the input of OSA from different fiber switch channels
Comb fiber switches
This is in the top section (High Level Control) | YJ_OR_HK |  |  |  |  | Select whether YJ or HK comb are selected .
Flattener Control and Status | [5 empty] | All of these are TBD
[6 empty] | Target counts | Target intensity (counts) observed on comb spectrometer of flattened level. This is controlled by Comb attenuation in "LFC Control (simple)". Not a separate control
Redundant | LFC_REFLT |  |  |  |  | Flatten | TBD whether to reflatten on command - see "Reflatten Comb (single event)"
Redundant | [5 empty] | Continuous | User indicates whether comb reflattening should occur continuously at intervals specified in ”update interval”. Not clear whether this will be an option or default. How do we (or would we want to) turn off flattener but still transmit comb. If we do, would we still want the comb light path to go through the flattener (i.e. turn off the voltage bias "map" on the SLM?)
Redundant | LFC_FLT_RTE |  |  |  |  | Update interval (seconds) | Time between flattening updates. The is the refresh rate, "Reflattener refresh rate" in "LFC Control (simple)", and may not be tunable if Menlo flattener is purchased.
| LFC_FLT_STAT |  |  |  |  | Flattener status | (TBD)

## new LFC function list
| Keyword | Value | Required Authorization to Write to Keyword | Read/Write | Type | Description | Additional Notes | Unit | Acceptable Range |  | Tolerance | Default Value | time period(unit:s) | Comb Python Module Function Name | Implementation status |  |  | Keyword description
[9 empty] | Min | Max |  |  |  |  | implemented in xml | read tested | write tested | brief | verbose
LFC High Level Control (Simple)
| LFC_RFOSCI_ONOFF | 1 or 0 | Y | R/W | boolean | set on or off of rfosc powersup | [6 empty] | 120 |  | complete | can test
| LFC_RFAMP_ONOFF | 1 or 0 | Y | R/W | boolean | set on or off of rfamp powersup | [6 empty] | 120 |  | complete | can test
| LFC_PTAMP_LATCH | none or 1 |  | R/W | boolean | reset arduino | value==none:get voltage//value==1:reset latch | [5 empty] | 120 |  | complete | TBD | TBD
| LFC_YJ_SHUTTER | 1 or 0 |  | W | boolean | set YJ shutter 1 or 0 | 1:pass light/0:shut light | [5 empty] | 120 |  | complete | can test | can test
| LFC_2BY2_SWITCH | 1 or 2 or none |  | R/W | enumerate | set 2 by 2 switch | 1:YJ/2:HK //none:get status |  | 1 | 2 |  |  | 120 |  | complete | complete | complete
| LFC_HK_SHUTTER | 1 or 0 or none |  | R/W | boolean | set hk shutter | 1 or open:pass light/0 or close:shut light | [5 empty] | 120 |  | complete | can test | can test
| LFC_VOA1550_ATTEN | db number |  | R/W | float(%.2f) | set attenuation db | 0-60db |  | 0 | 60 |  |  | 120 |  | complete | complete | complete
| LFC_VOA1310_ATTEN | db number |  | R/W | float(%.2f) | set attenuation db | 0-60db |  | 0 | 60 |  |  | 120 |  | complete | can test | can test
| LFC_VOA2000_ATTEN | db number |  | R/W | float(%.2f) | set attenuation db | 0-60db |  | 0 | 60 |  |  | 120 |  | complete | complete | complete
| LFC_IM_LOCK_MODE | 1 or 0 |  | R/W | boolean | set lock mode//read after write | value==1:pid mode//value==0:manual mode | [5 empty] | 120 |  | complete | can test | test after comb on
| LFC_EDFA27_INPUT_POWER | none |  | R |  | get input power of edfa27 | [6 empty] | 120
| LFC_EDFA23_INPUT_POWER | none |  | R |  | get input power of edfa23 | [6 empty] | 120
| LFC_MINICOMB_AUTO_SETUP | 1 or none |  | W |  | auto setup minicomb, im lock, if wrong, send email
| LFC_TEMP_MONITOR | none |  | R | boolean | auto temp monitor | if temp goes above threshold, trigger LFC_CLOSE_ALL and LFC_EMAIL | [7 empty] | complete | can test
| LFC_CLOSE_ALL | 1 or none |  | W | boolean | shut up all power device | value==1:shut up all output | [7 empty] | complete
LFC_EMAIL
| LFC_TEMP_TEST1 |  |  | R | float array(%.2f) | get all temp on USBdaq1 | test | [5 empty] | 20 |  | complete | complete
| LFC_TEMP_TEST2 |  |  | R | float array(%.2f) | get all temp on USBdaq2 | test | [5 empty] | 20 |  | complete | complete
| LFC_T_GLY_RACK_IN |  |  | R | float(%.2f) | [7 empty] | 20 |  | complete | can test
| LFC_T_GLY_RACK_OUT |  |  | R | float(%.2f) | [7 empty] | 20 |  | complete | can test
| LFC_T_EOCB_IN |  |  | R | float(%.2f) | [7 empty] | 20 |  | complete | can test
| LFC_T_EOCB_OUT |  |  | R | float(%.2f) | [7 empty] | 20 |  | complete | can test
| LFC_T_RACK_TOP |  |  | R | float(%.2f) | [7 empty] | 20 |  | complete | can test
| LFC_T_RACK_MID |  |  | R | float(%.2f) | [7 empty] | 20 |  | complete | can test
| LFC_T_RACK_BOT |  |  | R | float(%.2f) | [9 empty] | complete | can test
| LFC_RIO_T | lockout keyword | Y (lockout) | R/W |  | RIO pump laser temperature | Typically control temperature, monitor current | C | 10 | 25 |  | 19.181 | 20 |  | complete | complete | can test (10-25)
| LFC_RIO_I | lockout keyword | Y | R/W |  | RIO pump laser current | Typically control temperature, monitor current | mA | 145 | 155 |  | 150 | 120 |  | complete | can test | can test (145-155)
| LFC_EDFA27_P | lockout keyword | Y | R/W | float | Small EDFA (500 mW) 1 output power | Use constant power mode (APC mode) ; Read is monitored, write is setpoint. Value='default':apc mode 450mw | mW | 0 | 630 |  | 630 | 120 |  | complete | can test
| LFC_EDFA27_P_DEFAULT |  |  | R/W | float | Small EDFA (500 mW) 1 output power | value==none:read power/value!=none,set 'apc'mode, 450mw | mw |  |  |  |  | 120
| LFC_EDFA27_ONOFF | lockout keyword | Y | R/W | Boolean | Small EDFA (500 mW) 1 emission on/off | [6 empty] | 120 |  | complete | can test
| LFC_EDFA27_AUTO_ON | none or 1 |  | R/W | float | Small EDFA (500 mW) 1 auto turn on | value==none:read power/value !=none, set apc mode 450mw and turn on edfa27 | [5 empty] | 120
| LFC_EDFA13_P | lockout keyword | Y | R/W |  | Small EDFA (20 mW) 2 output power | Use constant power mode (APC mode) ; Read is monitored, write is setpoint | mW | 0 | 20 (TBD) |  | 20 | 120 |  | complete | can test
| LFC_EDFA13_ONOFF | lockout keyword | Y | R/W | Boolean | Small EDFA (20 mW) 2 emission on/off | [6 empty] | 120 |  | complete | can test
| LFC_EDFA23_P | lockout keyword | Y | R/W | float | Small EDFA (200 mW) 3 output power | Use constant power mode (APC mode) ; Read is monitored, write is setpoint. Value='default':acc mode 80mw | mW | 0 | 20 (TBD) |  | 20 | 120 |  | complete | can test
| LFC_EDFA23_P_DEFAULT |  |  | R/W | float | Small EDFA (200 mW) 3 output power | value==none:read power/value!=none,set 'acc'mode, 80mA | [5 empty] | 120
| LFC_EDFA23_ONOFF | lockout keyword | Y | R/W | Boolean | Small EDFA (200 mW) 3 emission on/off | [6 empty] | 120 |  | complete | can test
| LFC_EDFA23_AUTO_ON | none or 1 |  |  | float | Small EDFA (200 mW) 3 auto trun on | value==none, read emission status/value!=none,set acc mode 80mw and turn on power | [5 empty] | 120
| LFC_RFAMP_I | lockout keyword | Y | R/W |  | RF Amp current | default set:I=5A | Amp |  |  | +/- 0.2 | 4.4(Typical) | 120 |  | complete | can test | can test if off (check LFC_RFAMP_ONOFF, make sure it's off)
| LFC_RFAMP_V | lockout keyword | Y | R/W |  | RF Amp voltage | This is the control value; Caution! no voltage ramp allowed - can only go directly from off to 30 V.default set:30v | V | 30 | 30 |  | 30 | 120 |  | complete | can test | can test if off (check LFC_RFAMP_ONOFF, make sure it's off)
| LFC_RFAMP_DEFAULT |  |  | R/W |  | read/set RF amp default value | value==none, read voltage and current/value!=none,set v1=30v I1=5A | [5 empty] | 120
| LFC_RFAMP_ONOFF | lockout keyword | Y | R/W |  | set RF Amp on/off state | [6 empty] | 120 |  |  | can test
| LFC_RFOSCI_I | lockout keyword | Y | R/W |  | RF Oscillator input current | default set:I=3A | Amp | 0.35 | 0.7 |  | 0.41 (operating)/0.62(warming up) | 120 |  | complete | can test | can test if off (check LFC_RFOSCI_ONOFF, make sure it's off)
| LFC_RFOSCI_V | lockout keyword | Y | R/W |  | RF Oscillator input voltage | default:15v | V | 15 | 15 |  | 15 | 120 |  | complete | can test | can test if off (check LFC_RFOSCI_ONOFF, make sure it's off)
| LFC_RFOSCI_DEFAULT |  |  | R/W |  | read RF Oscillator v and I/set default v and I | Value==none, read V and I/value!=none, set V=15v I =3A | [5 empty] | 120
| LFC_RFOSCI_ONOFF | lockout keyword | Y | R/W |  | set RF osci on/off state | [6 empty] | 120 |  |  | can test
| LFC_IM_BIAS | lockout keyword | Y | R/W | float | Mini-comb amplitude lock offset bias | Amplitude lock voltage bias, is set by the locking loop, applied by RP FPGA | V | -3 | 3 | 0.001 |  | 120 |  | complete | can test
| LFC_IM_RF_ATT | lockout keyword | Y | R/W | float | Mini-comb RF voltage controlled variable attenuation setting | default:0.72v | V | 0 | 10 |  |  | 120 |  | complete | can test
| LFC_WSP_PHASE(TBD) | lockout keyword | Y | R/W | float | Waveshaper Dispersion Compensation | Controls the d2 dispersion setting in the Finisar Waveshaper | ps |  |  |  | 2.1 ps/nm | 120 |  | TBD | TBD
| LFC_WSP_ATTEN | lockout keyword | Y | R/W | float array shape=(2, N) | [7 empty] | 120 |  | TBD | TBD
| LFC_PTAMP_PRE_P | lockout keyword | Y | R/W | float | High-power EDFA pre-amp output power | monitors output power from the Pritel pre-amp. Default:600ma | ma |  |  |  | 135 | 120 |  | complete | can test
| LFC_PTAMP_PRE_P_DEFAULT |  |  |  |  | High-power EDFA pre-amp current | value=none, read pre amp current/value!=none, set to 600mA | [5 empty] | 120
| LFC_PTAMP_OUT | lockout keyword | Y | R |  | High-power EDFA output power | Read only, not set | W | 0 | 4 |  |  | 120 |  | complete | can test
| LFC_PTAMP_I | lockout keyword | Y | R/W |  | High_power EDFA output current set | default:3.8A | Amp | 0 | 4.2 |  | 3.6 | 120 |  | complete | can test
| LFC_PTAMP_I_DEFAULT |  |  | R/W |  | High Power EDFA output Power/current set | value=none, read out amp current/value!=none, set to 4.1A | Amp |  |  |  |  | 120
| LFC_PTAMP_ONOFF | lockout keyword | Y | R/W | Boolean | High-power EDFA emission on/standby | [6 empty] | 120 |  | complete | can test
| LFC_WGD_T | lockout keyword | Y | R/W | float | Waveguide TEC control |  | C | 22.5 | 23.5 |  | 23 | 20 |  | complete | can test | can test (+/- 0.5 degrees)
| LFC_PPLN_T | lockout keyword | Y | R/W | float | PPLN TEC control | [5 empty] | 37 | 20 |  | complete | can test | can test (+/- 0.5 degrees from where it is)
| LFC_HK_ATT |  | N | R/W | float | Comb Attenuation H and K band (Intensity) | VOA2000 db set. 0-60 dB attenuation. Default:60db | dB | 0.01 | 60 | 0.01 | 60 | 120
| LFC_FLT_ATT |  | N | R/W | Integer (? Need check) | Control position of the slider holding the neutral density filters in the spectral flattener | can't control now | dB | TBD | TBD | TBD | TBD | 120
| LFC_PMP_ATT |  | N | R/W | float | Pump Attenuation (Intensity) | voa1550 db set. Default:60db | dB | 0.01 | 60 | 0.01 | 60 | 120 |  | complete
| LFC_YJ_ONOFF |  | N | R/W | Boolean | Open or close the shutter of flattened YJ comb output | same as LFC_YJ_SHUTTER |  | TTL |  |  | 0 | 120
| LFC_HK_ONOFF |  | N | R/W | Boolean | Open or close the shutter after the HK flattener output | same as LFC_HK_SHUTTER |  | TTL |  |  | 0 | 120
| LFC_YJ_ATT |  | N | R/W | float | Comb Attenuation Y and J band (Intensity) coming from KPF comb (Nem) | voa1310 db set. Default:60db | dB | 0.01 | 60 | 0.01 | 60 | 120
| LFC_YJ_HK |  | N | R/W | String | Selects either HK or yJ comb to be sent to specctrograph | same as LFC_2BY2_SWITCH. Input value:1(represent YJ) or 2(represent HK) (input number only) | [5 empty] | 120
| LFC_RB_AUTO_LOCK | TBD
| LFC_IM_AUTO_LOCK | TBD
| LFC_WSP_OPTIMIZE | TBD
Note on U1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Please fill these in (used in help message for users, see LFC.xml.sin file)
Note on B7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Split into two keywords? Reply: How to read after reset?
Note on C7: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implemented as 1 or 0, 1=on and 0=off
Note on F9: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: I implemented this as enumerated
Note on F11: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Please provide formats for float types. For example, float(%.2f)
Note on C14: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: If this keyword value is going to be either “pid” or “man”, it could be implemented as enumerated type, as in LFC_2BY2_SWITCH Reply: Implementing as boolean
Note on B19: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Fault or okay - boolean Reply: 0: Fault, 1: Normal
Note on H19: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 0:Fault, 1:Normal
Note on N19: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Please fill in
Note on H20: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 0: running, 1: shut up all output
Note on B21: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Doesn’t have to be KTL keyword
Note on B23: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Maybe these are okay to remove? Reply: Keep them and change names
Note on F35: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implementing as float
Note on B36: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Not exposed to end user? Reply: Not needed
Note on F36: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Not sure what type it’s supposed to be
Note on B38: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Not needed
Note on F38: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Not sure what type it’s supposed to be
Note on B42: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Not needed
Note on B44: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Not needed
Note on B47: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Not needed
Note on B48: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Duplicate
Note on B51: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Not needed
Note on B52: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Duplicate
Note on F55: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Only 1-dimensional array allowed
Note on F57: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Implemented as float
Note on B58: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Not needed
Note on B61: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Not needed
Note on B69: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Duplicate
Note on B71: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Duplicate
Note on B72: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Duplicate
Note on B73: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Duplicate
Note on B74: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Duplicate
Note on B75: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Duplicate
Note on H75: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Is this needed, separate from LFC_2BY2_SWITCH?

## KPF Keywords
KPF LFC Control |  |  |  |  | Just listed to compare KPF Menlo comb functionality with HSF comb
| TBD |  | R/W | LFC On/Off
KPF Menlo Flattener (NOTE subject to change; current list is from NEID)
| LFCPCFSN | VIS0002 | R | LFC PCF Serial Number | N/A
| LFCPCFRT | 1590962+ | R | LFC PCF Runtime (s) | N/A
| LFCFLIR | 1332 | R | LFC Fiberlock IR Intensity | May be analagous to beat note SNR?
| LFCFLVIS | -36 | R | LFC Fiberlock VIS Intensity | N/A
| LFCF0 | 4954999997.833 | R | LFC Filtered Offset Freq (Hz) | For HSF, measure of beatnote between fiber comb tooth and RIO pump laser for EO comb?
| LFCF0REF | 4955000000 | R | LFC Filtered Offset Freq Ref (Hz) | Not sure if HSF analog?
| LFCFR | 20000000000 | R | LFC Filtered Rep Rate (Hz) | Same as HSF comb drive oscillator frequency reported on the Pendulum frequency counter; for fiber comb, would be a measure of the fiber comb rep rate (~180 MHz). See "Fiber Comb Health and Safety" spreadsheet.
| LFCFRREF | 20000000000 | R | LFC Filtered Rep Rate Ref (Hz) | Not sure if HSF analog?
| LFCSTAT |  | R | LFC Status Bitfield | Same as KPF LFC control status: on/off? For HSF, same as "State Reporting"

## Old list from PARVI
Number | Parameter | Type | Sent to | Received from |  | Description
DATA I/O
Commands to Control Comb
1 | Comb Attenuation(Intensity) |  |  |  | KPIC computer
2 | Pump Attenuation (Intensity) |  |  |  | KPIC computer
3 | Comb on |  |  |  | KPIC computer
4 | Comb standby |  |  |  | KPIC computer
5 | Comb off |  |  |  | KPIC computer
6 | Reflatten comb | [5 empty] | TBD – will be used if have SLM for spectral flattening
Comb Operating Parameter Health and State Reporting
7 | GPS locked | boolean | Comb computer
8 | Repetition Rate (comb spacing) | number | Comb computer
9 | Drive Oscillator Locked | boolean | Comb computer
10 | Comb Computer Time stamp(?) |  | Comb computer
11 | Pump center wavelength | number | Comb computer
12 | Pump frequency status (locked, standby) | text | Comb computer
13 | Comb board operating temperature (4 signals) | number | Comb computer
14 | State reporting (on, off, standby) | text | Comb computer
15 | Average counts | integer | Comb computer
16 | Comb spectrometer status | boolean | Comb computer
17 | Pump laser locked to reference | boolean | Comb computer |  |  |  | Pump line locked to tooth of fiber comb/beat note measure (will need separate reporting parameters for fiber comb performance)
18 | Filter cavity locked | boolean | Comb computer
Number | Parameter | Type | Sent to | Received from | Description
Advanced Comb Control Settings
18 | EDFA preamp I1 | number |  | KPIC computer
19 | EDFA preamp I2 | number |  | KPIC computer
20 | High Power EDFA pump current | number |  | KPIC computer
21 | Dispersion Compensation | number |  | KPIC computer
Comb Communications
22 | TCP Server |  |  | KPIC computer
23 | UDP Server |  |  | KPIC computer
24 | Server |  |  | KPIC computer
25 | Receiver |  |  | KPIC computer
26 | Last received command |  |  | KPIC computer
27 | Local Comb computer data collection |  |  | KPIC computer
28 | Data Directory |  |  | KPIC computer
29 | Update Period |  |  | KPIC computer
30 | OSA reporting time |  |  | KPIC computer
31 | Recording time |  |  | KPIC computer
32 | Record Start |  |  | KPIC computer
Flattener
33 | Target counts | integer |  | KPIC computer | Target intensity (counts) observed on comb spectrometer of flattened level.
34 | flatten | boolean |  | KPIC computer
35 | continuous | boolean |  | KPIC computer | User indicates whether comb reflattening should occur continuously at intervals specified in ”update interval”.
36 | Update interval (seconds) | integer |  | KPIC computer | Time between flattening updates

## 
Flattener Control and Status

## Test Case Selection
Phase 1 test
Query multiple times, 1Hz
| Purpose | Functionality | Keywords | Input | expected output | Process
| Check pump laser | Current of RIO | LFC_RIO_I
| Check pump laser | temperature of RIO | LFC_RIO_T
|  | RF Oscillator current | LFC_RFOSCI_I
|  | RF Oscillator voltage | LFC_RFOSCI_V
|  | RF amplifier current | LFC_RFAMP_I
|  | RF amplifier voltage | LFC_RFAMP_V
Amonics 27 power input
|  | Amonics 27 power output | LFC_EDFA27_P
Amonics 27 functioning mode (ACC/APC)
OSA trace
|  | IM Bias voltage | LFC_IM_BIAS
|  | YJ/HK fiber switches (Arduino) | LFC_YJ_OR_HK
YJ Shutter (Arduino)
|  | 1560 pump VOA (VOA1) | LFC_PMP_ATT
|  | Pritel preamp current | LFC_PTAMP_PRE_I
|  | Pritel pwramp current | LFC_PTAMP_I
|  | Pritel output power | LFC_PTAMP_OUT
Note on E16: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: (Yoo Jung) There is an "enumerated" keyword type that is integer numbered but also has ascii type representations. Users can use either integer number or string. (for example, 0: YJ, 1: HK), and if string is given it is translated to integer value. Might be easier for error handling because string inputs other than "YJ" and "HK" will be conveniently blocked by KTL.